## Slide 1: Why did i choose this subject
Mancardi Xavier (advisor : Patrice Verrecchia)
ENSICAEN – Université Paris-Sud
Subject covers every topic I approached during my studies : electronics, optical physics, nuclear physics, medical imaging
Instrumental thesis

## Slide 2: Thesis subject
Détecteur liquide multipixellisé pour l’imagerie médicale et préclinique
Conception and construction of the ionization part of the CaLIPSO detector for PET imaging modality : ultrapurification and caracterisation of the ionization parameters of the detection medium (TrimethylBismuth TMBi), electronic set up, Monte-Carlo simulation of the detector and performances analysis
Glossary : TrimethylBismuth, PET imaging modality, ultrapurification, electronics, simulation
CaLIPSO detector
PET principle

## Slide 3: Plan
Physical principle
  TEP
  CaLIPSO
  PET detectors : state of the art
Ultrapurification
  Why
  Adsorption
  Molecular sieves
  Impurities analysis
  Ultrapurification station
Tvie cell tests
  Tvie cell - charge mode
  Tvie cell – Gfi mode
  Electronic noise optimisation
Measurements
  Expected muon signal
  Gfi measurement
  γ 511 keV measurement
Future
  Multipixellised detector
  Capacitive readout
Conclusion

## Slide 4
I. PHYSICAL PRINCIPLE

## Slide 5: PET : Positron emission tomography
Radioactive tracer (ex : FDG) is injected in the patient body and then chemically bounded in tissue
β+ decay
Annihilation with an electron of tissue  2 511 keV γ are emitted back-to-back
2 γ detected in coincidence
Offline analysis and image reconstruction
Functional imaging technique in nuclear medicine
  Oncology : small tumors and metastases imaging
  Neurology : exams of neurodegenerative diseases (Alzheimer, Parkinson)
Important characteristics of PET detector
  Efficiency
  Spatial resolution
  Time resolution
PET principle

## Slide 6: CaLIPSO
Bi
CH
3
CH
3
CH
3
Organometallic liquid apolar and aprotic
Radiation length ~ 3 cm
Bi : Z = 83 highest Z non radioactive element
 Photoelectric fraction RPE = 47%
EPE = 420 keV (K shell/ relativistic in TMBi)
 Cerenkov emission (20 photons at 511 keV)
 Cerenkov light detection (trigger / z positioning)
 Optical detector operational
Drift of produced charges by the interaction of 511 keV γ
 charge detection (calorimetry and xy positioning)
 Ionization detector on study
Expected performances:
   3D positioning, 1 mm3
   time resolution ~ 100 ps
But : Double detection of the 511 keV γ interaction in the TMBi
TrimethylBismuth
CaLIPSO detector

## Slide 7: PET detector : state of the art
LSO/LYSO : The reference detector
LaBr3 : Excellent timing, poor PE efficiency, fair positioning
  Only relevant for full body, Time of flight PET config
  CdTe/CZT : Excellent position recontruction, poor PE efficiency
  Only relevant for single mouse PET imaging
CaLIPSO* : Best PE efficiency, Excellent positioning, very good timing
 Take the best of all technologies – Needed for high-res efficient Brain PET
* Inspired by the work done on liquid Xenon
Properties Detector | Atten Length (cm) | Coinc . PhotElecE Eff.( %) | Timing Resolution (ps, FWHM) | Energy Resolution (% FWHM) | G Interac. Postion. (mm) | End user friendly
LSO/LYSO | 1.23 | 12 | 300 - 500 | 10 | 2 to 10 | YES
LaBr3 | 2.3 | 1.9 | 100 - 300 | 3 | 4 to 10 | YES
CdTe/CZT | 2.0 | 2.2 | slow | 1- 3 | 0.1 | YES
CaLIPSO | 2.9 | 22 | ?150? - 380 | 10 | 0.15 | Will be !

## Slide 8
II. ULTRAPURIFICATION

## Slide 9: Why?
511 keV γ  4500 electrons
Electronegative impurities contained in the liquid reduce the signal
 ultrapurification needed
Commercial product obtained with a purity of 99,9995%
 chemically pure but insufficient for our application
 level of impurity required < 0,1 ppm oxygen equivalent
 lifetime for the electrons > 10 µs
Efficient technique for such level of purification
 molecular sieve

## Slide 10: adsorption
Selective retention of one or several compounds of a mixture on a solid surface
Adsorption of molecules < micropore size
Most polar molecules adsorbed first
Adsorption favorised by high pressures and low temperatures
TMB +
Impurities (Adsorbate)
TMB
Adsorbant
Zéolite 4A
Adsorption
physical
chemical
Electrostatic forces
Van der Waals forces
Chemical bond created

## Slide 11: Molecular sieves
Molecular sieve | Zeolite A | Zeolite X | Silica gel | Activated alumina
Pore size | 3-5 Å | 10 Å | 20-200 Å | 20-500 Å
Structure |  |  |  | 
4A zeolite structure
Metal aluminosilicates crystals
Product : monocrystals powder (particles ~ 10 µm)
Ionic structure
Negative charge Al- neutralised par Na+
Cation determinates pore size
Pellets
Powder
Pore size distribution of several molecular sieves

## Slide 12: Impurities analysis
Chart: 
| 2 | 12 | 13 | 14 | 15 | 16 | 26 | 27 | 28 | 29 | 36 | 50 | 108 | 123 | 182 | 209 | 210 | 221 | 222 | 223 | 224 | 225 | 238 | 239 | 240 | 254 | 255
 | 0.15147 | 0.11958 | 0.36671 | 1.4031 | 5.6441 | 0.31888 | 0.22321 | 0.62181 | 0.36671 | 0.2551 | 0.11161 | 0.071747 | 0.055804 | 0.063776 | 0.15944 | 100 | 2.551 | 0.9088 | 3.396 | 8.801 | 65.3061 | 1.2755 | 0.23119 | 75.5102 | 1.2755 | 40.3061 | 1.2117
In collaboration with DEN (V. Dauvois et D. Doizi)
Analysis of impurities contained in TMBi
Mass spectrum
Infrared spectrum
IR cell (left) et mass spectrometer cell (right)
TMBi chromatogram

## Slide 13: « DYNAMIc » ultrapurification station
TMBi in vapor phase passing on zeolites
Ultra high vacuum components (leak< 10-11 mbar l/s)
Passing through 4A zeolites in pellet form inefficient  powder
Passing through activated alumina (Al2O3)  reaction with impurities contained inside the sieve  game over for Al2O3
Dynamic station
Dynamic station plan

## Slide 14: « static » ultrapurification station
Soaking of TMBi in liquid phase in zeolites
13X zeolites, Silica gel  TMBi goes inside the pores and react with the walls eventually  game over for large pores in static station
3A zeolites powder form  first ionization signal
4A and 5A zeolites in powder form  on going
Static station
Static station plan

## Slide 15
Tvie cell electronic readout
Tvie cell
III. TVIE CELL TESTS

## Slide 16: Tvie Cell - Charge MODE
Optical Demonstrator optimized
E1
E2
Measure of ionization yield, free electrons lifetime and mobility
Low-noise iDeF-X electronic developed at SEDI (O. Gevin et al.)
Frisch grid needed for :
 signal to be independent of the interaction position in the detector
 Collecting only electron charges
Detector
Frisch grid
Tvie cell plan
  Charge impulsion readout
IDeF-X HD picture and layout

## Slide 17: Tvie cell – gfi mode
Test cell for Gfi measurement and high voltage behaviour checking
Tvie cell plan
Ionization current measurement
Ionisation chamber plan
E
  DC current readout
Gfi (number of free electrons created for 100 eV energy deposited)

## Slide 18: Electronic noise optimisation
Very weak signals (pA, fC)
Eliminate every noise sources
  Very low noise iDeF-X electronic
  Same ground for every electronic modules and detector
  High voltage filters
Charge preamplifier noise diagram
High voltage low-pass filter
White noise

## Slide 19
IV. MEASUREMENTS

## Slide 20: Expected muon signal
Detection in coincidence of a muon on two scintillators
µ
Energy deposit at minimum ionizing
Ionization signal readout with IDeF-X
Acquisition on Matacq card developed at SEDI (E.Delagnes et al.)
2,7 MeV average energy deposit (σ = 0,93 MeV)
Energie deposited by muons in TMBi in MeV
Matacq card
Chart: E(MeV)
| 100 | 200 | 250 | 300 | 350 | 400 | 500 | 600 | 700 | 800 | 900 | 1000
Méthode précise | 2.774 | 2.723 | 2.69 | 2.688 | 2.693 | 2.71 | 2.723 | 2.75 | 2.775 | 2.78 | 2.8 | 2.834
Energy loss regarding muon energy

## Slide 21: MUON signal (preliminary)
Drift time = 725 ns
With a Frisch grid the rise of the signal corresponds to the electron drift time between the grid and the anode

## Slide 22: GFI measurements
  Readout of the ionization current produced by a 60Co gamma rays source
Cobalt 60 source
γ
Current measurements regarding HV
With = ionization current
Δε = Edeposit*Asource*AcceptanceTVie
Examples of Gfi curves
TMBi/TMSi Gfi ratios

## Slide 23: Modelisation : Impact of liquid purity
TMBi brut (lifetime= 50 ns)
TMBi ultrapure (lifetime= 50 µs)
Energy deposited by Cobalt 60 in TMBi
Pure TMBi needed for calorimetry
Minimum noise needed to get as close as possible to the real spectrum

## Slide 24: ModElisation : 511 kev signal
γ 511 keV
γ 511 keV
  Detection in coincidence of two 511 keV γ from a β- 22Na source
Ultrapure TMBi required for calorimetry
Final test for ultrapurification validation

## Slide 25
V. FUTURE

## Slide 26: Multipixellised detector
In order to localized the interaction position on the transverse plan we need to pixellise the charge collector with 1mm² pixels (at best)
Number of pixels
Compact readout electronics needed
Pixel size effect on signal collection
Diaphony effect/ capacitive coupling between pixels
Pixellised detector

## Slide 27: Capacitive readout
In order to no longer need hermetic junctions for the signal readout on each pixel
 resistive layer (alumine 300 µm) and pad readout by capacitive influence*
 anode deposit tests on alumine performed by the CSNSM in Orsay (L. Dumoulin et L. Berge)
Capacitive readout principle
E2
*D.Attié et al., Piggyback resistive Micromegas , ArXiv 1208.6525
Tpeak < 10 µs

## Slide 28
VI. CONCLUSION

## Slide 29: CALIPSO detector perspectives
Operational
  Optical detector
  Ionization electronic readout
  Tvie cell modelisation
  First ionization signal after purification on 3A zeolites in powder form
On going
  TMBi ultrapurification
  Tvie cell performance study
  Mecanical conception of final CaLIPSO detector with the bureau d’étude
  Physical modelisation of final detector

## Slide 30
Thank you for your attention