## Slide 1: Retour d’expérience après test en faisceau à DESY
P. Colas

## Slide 2: Situation: on vient de tester la chaine complète avec 1 des 7 modules
New detector : new routing to adapt to new connectors, lower anode resistivity, new res. foil grounding on the edge of the PCB.
New 300 points flat connectors
New front end: keep naked AFTER chips and remove double diodes (count on RF to protect against sparks)
New Front End Mezzanine (FEMI)
New backend ready for up to 12 modules
New DAQ, 7-module ready and more compact format
New trigger discriminator and logic.

## Slide 3: Prochaines étapes
Production de 9 modules avec leur électronique
  Commande 9 PCB détecteurs au CERN?
  Appels d’offre FECs
  Réalisation des FEMIs
Banc test au CERN
  Etude thermique
  Homogénéité
  Gain
Préparation 7 modules (fibres, HV, LV, soft)

## Slide 4: Détecteur
Fonctionne bien. Resolution 60-75 microns. Neff 26-32
(excluant les lignes avec ASICs en mauvais contact).
A refaire à l’identique, avec davantage de controles et plus fortes tolérances mécaniques.
Chart: 
| 100 | 200 | 400 | 500 | 600 | 700 | 900 | 1000 | 1200 | 1500
 | 0.07209 | 0.075756 | 0.081603 | 0.084167 | 0.08791 | 0.086374 | 0.092065 | 0.087794 | 0.092663 | 0.09276

## Slide 5: Trigger
Fonctionne bien
Cables BNC-lemo commandés
Petites améliorations en cours

## Slide 6: DAQ
Vérifier le mapping
Visualisation des pédestaux moyens et rms indispensable
Suppression de zero avec fenetre glissante

## Slide 7: FECs
Retirer les résistances de protection, fignoler la CAO grace à la place gagnée.
Nouveaux appels d’offre, commandes: séparément PCB (CIBEL?) et cablage (?)
Stager la production

## Slide 8: FEMs
Produire 7 à 9 FEMIs identiques aux deux protos.

## Slide 9: Ancillaires
HV : utiliser celle d’EUDET? Boite de distri + N471A?
Fibres
LV . Alimentations? Cables?
Refroidissement par air

## Slide 10: Beam Test
Monday May 2 to Friday May 14
Up and running in 24 hours, but 2 days for debugging monitoring, mapping, tuning HV, etc…
All the initial program has been done (250 good runs of 5000 evts, + lots of cosmics).
Dismounting was very easy (3 connectors : HV, LV, fiber), module+FEelectronics in one piece.

## Slide 11: Beam size, Vdrift
Fitted rms 1.3 bins -> 3.9 mm
Vdrift checked OK

## Slide 12: Integrated electronics for 7 module project

## Slide 13: Integrated electronics for 7 module project
Remove packaging and protection diodes
Wire –bonding on board for AFTER chips
Use 2 × 300 pins connector
Replace resistors SMC 0603 by 0402 (1 mm × 0.5 mm)
25 cm
14 cm
0,78 cm
0,74 cm
4,5 cm
12,5 cm
2,8 cm
3,5 cm
3,5 cm
FEC
Chip

## Slide 14: Design of the integrated electronics

## Slide 15: Front-End Card (FEC)
4 chips per FEC wire bonding on the 8-layer PCB
Protections: a resistor (0201 SMC) at each channel input
  A0: 0Ω
  A1: 3Ω
  A2: 5,1Ω
  A3: 10Ω
Temperature measurement device for each FEC
Check after 2 weeks of operation: no ASIC lost
  The resistive foil protects against sparks

## Slide 16: First prototype of the electronics

## Slide 17
Thermal studies. IR camera shows hot spots (regulators, ADC)
2-phase CO2 cooling under study (KEK, Nikhef)

## Slide 18: Temperature during switching on
Temperature in the hall
Chart: 
| 2 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | 45 | 50 | 55 | 60 | 65 | 135
FEC 0 | 29 | 34 | 39 | 41 | 43 | 44 | 44 | 45 | 45 | 45 | 45 | 46 | 46 | 46 | 46
FEC 1 | 30 | 36 | 43 | 46 | 49 | 50 | 51 | 51 | 51 | 51 | 51 | 52 | 52 | 52 | 53
FEC 2 | 30 | 37 | 44 | 48 | 50 | 52 | 52 | 53 | 53 | 53 | 53 | 54 | 54 | 54 | 54
FEC 3 | 30 | 37 | 45 | 49 | 51 | 52 | 53 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 55
FEC 4 | 30 | 37 | 44 | 47 | 49 | 51 | 51 | 52 | 52 | 52 | 52 | 52 | 53 | 53 | 53
FEC 5 | 29 | 34 | 39 | 41 | 43 | 44 | 44 | 44 | 45 | 45 | 45 | 45 | 45 | 45 | 46
FEMI | 30 | 37 | 42 | 45 | 47 | 47 | 48 | 48 | 49 | 49 | 49 | 49 | 49 | 49 | 49

## Slide 19: Temperature after cooling
Chart: Temperature (°c)
| 1 | 5 | 10 | 15 | 20 | 25 | 30 | 31 | 33 | 35 | 40 | 45 | 46 | 48 | 50 | 55 | 60 | 70
FEC 0 | 45 | 43 | 42 | 41 | 41 | 40 | 40 | 40 | 39 | 39 | 38 | 38 | 38 | 38 | 38 | 37 | 37 | 37
FEC 1 | 52 | 49 | 47 | 46 | 46 | 46 | 45 | 45 | 44 | 43 | 43 | 42 | 42 | 42 | 41 | 41 | 41 | 41
FEC 2 | 54 | 52 | 50 | 48 | 48 | 48 | 48 | 47 | 47 | 46 | 45 | 45 | 44 | 44 | 44 | 43 | 43 | 43
FEC 3 | 54 | 53 | 51 | 49 | 49 | 49 | 48 | 48 | 47 | 47 | 46 | 45 | 45 | 44 | 44 | 44 | 43 | 43
FEC 4 | 53 | 51 | 49 | 48 | 48 | 47 | 47 | 46 | 46 | 45 | 44 | 44 | 43 | 43 | 43 | 42 | 42 | 42
FEC 5 | 46 | 45 | 44 | 43 | 42 | 42 | 42 | 41 | 41 | 40 | 40 | 39 | 39 | 38 | 38 | 38 | 38 | 38
FEMI | 47 | 44 | 42 | 41 | 41 | 41 | 40 | 40 | 39 | 38 | 37 | 37 | 37 | 36 | 36 | 36 | 35 | 35
Increasing of the Nitrogen flow
Temperature in the hall

## Slide 20: Event samples in B=1T