## Slide 1
L. An2, D. Attié1, Y. Chen2, P. Colas1, M. Riallot1, H. Shen2,
W. Wang1,2, X. Wang2, C. Zhang2, X. Zhang2, Y. Zhang2
2011 Nuclear Science Symposium
and Medical Imaging Conference
October 27th, 2011 – Valencia, Spain
R&D of a Fast-Neutron Imaging Detector Based on Bulk-Micromegas TPC
(1)
(2)
Notes: I’m here to present a TPC review

## Slide 2: Overview
Introduction: idea of Fast Neutron Imaging detector
Micromegas TPC for neutron imaging
Description of T2K electronics and the detector
Data analysis and results
Conclusion

## Slide 3: Characteristics and simulation of FNI detector
Characteristics expected of Fast Neutron Imaging detector based on TPC:
High spatial resolution: <100 µm
high quality imaging from Micro-Pattern Gas Detector
as Micro-Mesh Gaseous Structure (Micromegas)
  Low efficiency: ~ 0.01-1%,
  subject to thickness and kind of converter
  suitable for beam monitor/profile
  imaging in very high flux
Simulation tools:
  Geant4 (physics processes)
  Garfield (gas processes):
    ionization energy
    electron drift velocity
    electron avalanche

## Slide 4: Simulation by Geant4 + Garfield
Data reconstruction method:
  identify cluster (track)
  extract hit position where the time is maximum tmax  interaction point
  integrate all events  image
Neutron event interacting
with polyethylene foil
and knocking out a proton
Garfield
Avalanches
n
p
e- avalanche
Drift lines from primary ionization e-
Proton track
X-Y readout plan
Drift time
 = 91.9 µm
p
Avalanche drift time
y-z readout plane

## Slide 5: Geant4 simulation for converter efficiency
CH2
gas
n
n, p
1 cm
6 cm
10 cm
25 µm – 20 cm
Neutronproton scattering efficiency in a polyethylene [C2H4]n layer
coming from 241Am-9Be source
For 100 000 events in the neutron spectrum:
Incident neutron spectrum

## Slide 6: Micromegas TPC for neutron imaging
gas
128 µm
HVmesh
Eamp ~ 30 kV/cm
10 mm
HVdrift
Edrift ~ 200 V/cm
Wax
Pb
Detector layout: 1728 (36 ×48) pads of 1.75 mm × 1.50 mm
Gas mixture: Argon + 5% Isobutane
+ bulk Micromegas
Elastic scattering on hydrogen n  p
+ masks (Pb, paraffin wax)
PCB
Micromegas
n
p
Aluminized polyethylene 25 µm
between 2 layers (0.5 µm) of Al
57.4 mm
88.6 mm
Cosmics
(x, y, t)

## Slide 7: Description of T2K electronics
Electronics designed at CEA/Irfu for the T2K TPC
AFTER-based electronics (72 channels/chip):
  low-noise (700 e-) pre-amplifier-shaper
  100 ns to 2 µs tunable peaking time
  full wave sampling by SCA
  frequency tunable from 1 to 100 MHz (most data at 25 MHz)
  12 bit ADC (rms pedestals 4 to 6 channels)
  full-scale gain from 120 fC to 600 fC
  zero-suppression capability
6 Front-End Cards (FEC) read out by a
Front-End Mezzanine (FEM)
Trigger signal needed
Spark protection
FEC with 4 AFTER chips

## Slide 8: Detector + electronics setup
Trigger from
Micromegas signal
FEM
FEC
Shielding
Window for x-rays source

## Slide 9: Performances of the Micromegas detector
Gain curve measured from 5.9 keV line using 55Fe source. Signals read out on the mesh in Ar/Isobutane 5%: G~103 @ 300 V
Energy resolution of ~12 % due to detector capacitance and noise
best energy resolution measured for a bulk Micromegas (~7 %)
Operating gas gain < 1500 and electronics full-scale gain set  360 fC
in order to cut the gamma-rays and cosmics events
 = 12 %

## Slide 10: 241Am–9Be source
Located in Yuzhong (near Lanzhou city), data taking in July 2011
Intensity: ~6 ×106 Hz (4π)
Neutron energy spectrum, according to ISO 8529 (reference radiations for calibrating neutron-measuring devices)
Mean energy ~4.5 MeV, up to 11 MeV
Data sample from source
36
48
Energy (MeV)
Source strength

## Slide 11: Proton/gamma-ray discrimination
~ 20 cm of paraffin in front of the detector
Cluster size is maximum at ~4
Equivalent charge: Landau MPV at ~40 keV
Uniform time spectrum
 60 keV (241Am) +  from neutron ?
Cluster size
Cluster charge
Time spectrum

## Slide 12: Proton/gamma-ray discrimination
6 mm of Pb in front of 8 cm of paraffin
before detector
Smaller cluster size
Equivalent charge: peak at ~110 keV
+ continuum up to 1 MeV
Doublet in time structure
 neutron signature ?
Cluster size
Cluster charge
Time spectrum

## Slide 13: Imaging with Lanzhou mask
Thickness: 17 mm
3 mm
Pb
Paraffin
+
Imaging
Counting
mode
Tracking +cuts in time & charge

## Slide 14: Imaging with CEA mask
Counting
mode
Thickness: 17 mm
3 mm
Pb
Paraffin
Imaging
Tracking +cuts in time & charge
+

## Slide 15: Imaging using others masks
1.5 mm
3 mm
3.5 mm
5 mm
2.5 mm
Thickness: 17 mm

## Slide 16: Conclusion
Since July 2011, the detector is ready for neutron imaging data taking
Still need to optimize the converter and the drift space
Find and use a high flux of fast neutron beam (D-T source) to avoid gamma-ray from source
Proton/gamma-ray discrimination should be improved by taking data with better neutron and gamma-ray stoppers