## Slide 1
Recent Ingrid Studies and DESY Test-Beam
Results with Octopuce
A. Chaus, D. Attie, P. Colas, M. Titov
CEA Saclay, Irfu/SPP, France
Ingrid
OUTLINE:
“Octopuce” Uniformity Studies in the Laboratory
DESY Test-Beam Track Reconstruction with Octopuce
Tests of IZM-3 InGrids
Future InGrid Tests using Low Energy Electrons from PHIL at LAL
International Workshop on Future Linear Colliders University of Texas at Arlington, October 22-26, 2012

## Slide 2
Each pixel can be set to:
● TOT ≈ integrated charge
● TIME = Time between hit and shutter end
“InGrid” Concept:
Protection Layer (few mm)
50 - 80 mm
(e.g. Si3N4)
(e.g. SU8)
Bump bond pads for Si-pixel
Detectors - Timepix or
Medipix2 (256 × 256 pixels
of size 55 × 55 μm2) serve
as charge collection pads.
3D Gaseous Pixel Detector  2D (CMOS pixel chip readout) x 1D (drift time)
Integrated Electronics: Pixel Readout of Micro-Pattern Gas Detectors
Through POST-PROCESSING INTEGRATE
MICROMEGAS directly on top of CMOS chip

## Slide 3
“InGrid’ Technology and “Driving” Developments
2011: Major Step Forward 
InGrid Production on a wafer level (107 chips)
2005: Single “InGrid” Production
2009: “InGrid” Production on a
3 x 3 Timepix Matrix

## Slide 4
The “Octopuce” (2009)
HV connections
“Grid irragularities”
Non homogeneous area on the mesh of different chips
chip 1
chip 4
chip 5
chip 8

## Slide 5
5
Chip 1
Octopuce Studies & Fe55: Homogeneous Irradiation
(pixels)
(sum of
pixels TOT)
Fe55 Studies (26/07/2012):
He/Iso 80/20
Vmesh=390 V
Vdrift =3000 V
Integrated picture for all 8 chips:
(discontinuity = distance
between pixels in the cluster)
Chip 4
“Hot spots”
Chip 5
Chip 8

## Slide 6
Octopuce: Uniformity of the Response (I)
THRESHOLD = 1000 COUNTS
THRESHOLD = 2000 COUNTS
THRESHOLD = 4000 COUNTS
Fe55 Studies (26/07/2012):
He/Iso 80/20
Vmesh=390 V
Vdrift =3000 V
Non-sensitive (~ 1.5 mm)
areas between chips
Chip 1
Chip 4
Chip 5
Chip 8

## Slide 7
Octopuce: Uniformity of the Response (II)
Non-sensitive (~ 1.5 mm)
areas between chips:
THRESHOLD = 6000 COUNTS
THRESHOLD = 8000 COUNTS
THRESHOLD = 10000 COUNTS
More studies required to understand:
if “dead” areas between chips decreases
with increased drift field;
“dead” area on the outer edge as a
function of the guard voltage

## Slide 8
Expected number of primary electrons in He/Iso (80/20) ~ 165
Octopuce: Cluster Size Distribution
Single electron sensitivity is very high for all (but 1 and 4) chips

## Slide 9
Chip number | Threshold level
1 | 337
2 | 335
3 | 340
4 | 335
5 | 330
6 | 327
7 | 330
8 | 320
Octopuce: Total Cluster Charge
Some differences in amplification gaps between different chips are seen by microscope
(studies are not conclusive, might come from different thickness of the glue under the chips)
Chips 1 and 4 have a lower response (same trend as for the cluster size distribution) 
Difference in amplification gap
Difference in threshold (too big to explain differences between 1 &4 and others)
After equalization:

## Slide 10
DESY Test-Beam with Octopuce & Large Prototype TPC (December 2010)
B = 0 T
He/iC4H10(80:20)
B = 1T
He/iC4H10(80:20)
4 chips
2 chips

## Slide 11
Timepix pane for Large Prototype TPC
Mezzanine board
Mother board
Guard ring
Chips on a mezzanine board making wire bonding easier
Large Prototype compatible
Heat dissipator
“The Octopuce” in the Large Prototype TPC

## Slide 12
Octopuce in the Large Prototype TPC at 0 T
He/Iso 80/20
Vmesh = 380V
Time mode
Shutter Time: 100 ms, start given by beam trigger
Occupancy is very low:
< 0.1 % (no noisy frames)
Example of a typical event:

## Slide 13
ROOT based analysis package developed by John Idarraga (LAL)
C++ classes including processors:
  OctoCEA (define in a few minutes)
  Pattern recognition of tracks for low threshold
MAFalda: Medipix Analysis Framework
Use MAFalda as the first step  implement reconstruction algorithm in Marlin TPC

## Slide 14
MAFalda: Cluster Finding and Track Segments (I)
Form cluster from the pixels with “discontinuity” < 40 pixels (each cluster
should contain > 12 pixels)
Calculate the linear regression with all points in the cluster (red dotted line)
Calculate residuals from the red line to each point (if > 80 % the points are within 20 pixels – this is the “track segment” !

## Slide 15
MAFalda: Cluster Finding and Track Segments (II)
Check incident angles of track segments
 Take segments within +/-10 degrees
(electron beam is in the horizontal direction)
Number of segments
per reconstructed track
(most tracks consists of
2-3 segments):

## Slide 16
Track Reconstruction
Y = ax +b
b : 417.558 | a : -0.021984
b : 421.54 | a : -0.0199386
b : 428.789 | a : -0.0341787
b : 431.063 | a : -0.0362019
Reconstruct track from track segments (if more than 2 segments):
Apply data quality cuts to make sure all segments correspond to a given track
Perform a linear regression for all pixels on a track  reconstructed track (blue dotted line)

## Slide 17
Track Reconstruction: Residuals
An unbiased estimate of the single point resolution: s = Sqrt(s1 * s2):
Perform fit to all pixels and calculate the distance between each pixel and the position
of the point of closest approach along the fitted track:
s1 = Gaussian fit, when all pixels are included in the track fit
s2 = Gaussian fit, omitting pixel under consideration from the track fit
in number
of pixels

## Slide 18
DESY Test-Beam: Track Residuals
Summary of residuals (s = sqrt(s1*s2)) for different Z-coordinates:
(corresponds to the different positions of electron beam
passing through the Large Prototype TPC)

## Slide 19
DESY Test-Beam: Track Residuals as a function of Z
(Large Prototype TPC)
PRELIMINARY:
PRELIMINARY:
Very low statistics, behavior needs to be understood

## Slide 20
Two micro-TPC boxes have been built
Drift distance in micro-TPC (~ 10 cm)
is large enough to allow study of single
electron response from Fe55 source
Saclay Micro-TPC with Timepix
TIMEPIX
Micro-TPC
Fe55

## Slide 21
Studies of new IZM-3 InGrids in the Saclay micro-TPC
2011: Major Step Forward  InGrid Production on a wafer level
2013: 3rd IZM production run to post-process Timepix chips on a wafer level
Received 6 IZM-3 InGrids in Saclay (earlier studies
with IZM-3 InGrids have been performed in Bonn,
NIKHEF)
Four chips are mounted on PCB (one does not work)
One InGrid is tested 
in general, good behavior
Guard ring problem
in micro-TPC
Some noise and/or
discharges
Some local grid
issues

## Slide 22
Proposal of a Flexible Detector Setup
using Low Energy Electrons from PHIL at LAL
Joint proposal LAL & IRFU
LAL contribution from S. Barsuk, L. Burmistrov, H. Monard, A. Variola
Goal: obtain samples of “monochromatic” electrons
with energy between 1 and 5 MeV and energy spread of better than 10%
with adjustable intensity down to 104 electrons per bunch
PHIL provides electrons with momentum 5 MeV/c and 109 particles per bunch
Study dE/dx by cluster counting using InGrid detectors
the electron range 1-5 MeV
(earlier simulation results by M. Hauschild & NIKHEF
experimental studies)

## Slide 23
Setup idea:
Use electrons from PHIL
Reduce energy/intensity using Al plug
Select unique direction for electrons passing the plug with collimator 1
Select required energy by half-turn of electron in the magnetic field (position of collimator 2)
Adjust intensity/energy spread using collimator 2, positioned in front of tested detector
Spectrometer to sample “monochromatic” low energy electrons

## Slide 24
Momentum and angular spectra of electrons passing through the Al plug, depending
on the plug thickness : Geant4 simulation
5 MeV/c electron
q
Aluminum plug