## Slide 1
Physical principle
Ultrapurification
Single pixel detector
Electronics
Measurements
Conclusion
Technologie détecteur de charge ultravide multipixellisé avec électronique ASIC bas bruit et ultrapurification
Xavier Mancardi

## Slide 2: Physical principle
Ionisation
Cherenkov Emission
Electrons drift
Time resolution
Z position*
XY position, Z position*
Energy
Optical detector
Ionization detector
Expected performances:
   3D positioning: 1 mm3
   time resolution ~ 100 ps
   energy resolution ~ 10%
CaLIPSO
PET principle
CaLIPSO detector principle
Xavier Mancardi

## Slide 3: ultrapurification
Electronegative impurities reduce the signal by capturing electrons
 ultrapurification < 0,1 ppm oxygen equivalent (electrons lifetime > 10 µs)
 molecular sieve (3A, 4A, 5A and 13X zeolites)
 batch and column devices
Batch (r.) and column (u.) ultrapurification stations
Xavier Mancardi

## Slide 4: Single Pixel Detector
Ionization chamber
Very weak signals (pA, fC)  Eliminate every noise sources
Charge collection surface
Frisch grid
Optical Demonstrator
E1
E2
Titanium layer on sapphire
Single pixel cell picture
Single pixel cell
Single pixel cell electronic readout
Xavier Mancardi

## Slide 5: Electronics
<
IDef-X chip
Charge pre-amplifier  Low-noise iDeF-X electronic developed at SEDI (O. Gevin et al.)
  Equivalent Noise charge (ENC) = 33 e- + 5,5 e-/pF
  Gain max = 200 mV/fC
  32 independent channels (ASIC)
Schematic of an IDeF-X channel
Test environment
Xavier Mancardi

## Slide 6: Measurements: 3 methods
Detection in coincidence of a muon on two scintillators : Electron mobility
µ
2γ 511 keV
Cobalt 60 source
γ
Chart: I = f(E)
| 0 | 16 | 100 | 200 | 350 | 500 | 750 | 1000 | 1500 | 2000 | 2500 | 3000 | 3500 | 4000
 | 34 | 40 | 106.6667 | 148 | 163.6667 | 172 | 177 | 175.6667 | 175.6667 | 185 | 186.6667 | 198.3333 | 200 | 193.3333
  Current produced by a 60Co gamma rays source : Charge production yield (Gfi)
  Detection in coincidence of two 511 keV γ from a β+ 22Na source : Ultrapurification
Xavier Mancardi

## Slide 7: conclusion
Now Operational
  Optical detector
  Ionization electronic readout
  Single pixel cell modelisation
  First ionization signals after purification on 3A zeolites in powder form
On going
  TMBi ultrapurification
  Single pixel cell performance study
  Mechanical conception of final CaLIPSO detector with « bureau d’étude »
  Physical modelisation of final detector
Xavier Mancardi

## Slide 8
Thank you for your attention
Xavier Mancardi

## Slide 9: Future : multipixelisation – UHV Vacuum
No feedthrough 
Alumine 300 µm + resistive layer => readout by capacitive coupling
Thin film deposits on alumine performed by the CSNSM in Orsay (L. Dumoulin et L. Berge)
Signal readout by capacitive coupling
E1
E2
Pixellised detector
Capacitive readout principle
Xavier Mancardi