## Slide 1: Optical readout: an intuitive readout for MPGD detectors
Filippo Resnati (EP-DT-DD) on behalf of the GDD team
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 2: Overview
Relevant mechanisms for the operation of gaseous detectors
Imaging applications of MPGDs outside high energy physics
Developments of the optical readout for MPGDs at CERN
Possible other applications
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 3: Overview
Relevant mechanisms for the operation of gaseous detectors
Imaging applications of MPGDs outside high energy physics
Developments of the optical readout for MPGDs at CERN
Possible other applications
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 4: Why gaseous detectors
Large dimensions (several m2)
Not expensive (< 5 kCHF/m2)
Good position resolution (<100 um)
Excellent time resolution (<5 ns and even <100 ps)
Radiation hard (10s year of LHC)
Low material budget (<0.01 X0)
Compatible with magnetic field
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 5: Gas as active medium
Gas at RTP is not dense
Small amount of energy released
Few primary electron-ion pairs
Amplification is needed
Two possibilities: charge and light
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 6: Townsend mechanism
Externally applied electric field acts on electrons and ions (free to move) separating them.
On top of the random motion (diffusion) charges acquire energy and move following the field lines (drift).
If the field is high enough, the energy gained by the electrons is sufficient to further ionise the gas.
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 7: Townsend avalanche
+
-
+
-
Cathode
Anode
For a linear response of the detector, the amplification must be confined in a small region.
α: average number of electron ion pairs created by a single electron per unit path
Avalanche increases ~exponentially in time, field and space
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 8: MWPC
The field is shaped by the curvature of the wire.
Several wires to allow the position resolution.
The field is sensitive to wire imperfections.
Slowly moving ions affect the rate capabilities.
E
r
avalanche
drift
10s μm
+
-
+
-
-
-
Cathode
Anode
10s μm
Charpak, 1968
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 9: Change of paradigm (I)
In order to evacuate the ionic charge fast, the cathode electrodes are placed in the vicinity of the anode electrodes:
Micro Pattern Gaseous Detectors.
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 10: MSGC
Ions are evacuated fast: better rate capability.
Finer anode pitch is possible: better position resolution.
The dielectric substrate charges up, changing the field.
Close anode and cathode occasionally result in sparks.
+
-
-
Anode
Cathode1
Cathode2
10s μm
Substrate
Oed, 1988
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 11: Change of paradigm (II)
In order to reduce the effects of the discharges and the ageing, the electric field responsible for the amplification is shaped by extended electrodes.
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 12: MPGDs: a rich variety
micro
electrode
hole
mesh
MicroPin
MicroDot
MicroPic
MHSP
MicroGroove
MicroWell
Micro-bulk MM
RPWell
Micro-RWell
GEM
TH-GEM
LEM
GEMpix
MicroGap
MicroWire
MSGC
Non-exhaustive list
extended
type
MicroMeGas
CAT
InGrid
GridPix
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 13: MicroMeGas
+
-
-
Anode can be made resistive to reduce the discharge effects
Cathode
Segmented anode
Pilar
~130 um
Metal mesh
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017
Giomataris, 1996

## Slide 14: GEM
+
-
-
Anode1
Cathode1
Anode2
Cathode2
70 μm
50 μm
Amplified electrons available for further amplification
-
-
-
Sauli, 1997
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 15: GEMs
-
-
-
-
-
-
-
-
-
-
-
+
-
-
-
-
Push further the discharge limit: larger gain
-
-
-
-
-
-
-
-
-
-
Sauli, 1997
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 16: Scintillation
Excitation happens at lower fields. It results in scintillation.
Electron and ion recombination may result in scintillation too.
total
Excitation threshold
Ionisation threshold
Electron argon cross sections
Electron energy distribution
Magboltz computation
Electrons in pure argon
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 17: Scintillation
Valid for all the rare gases
Simplified, but effective picture
Unlikely because of what follows
First continuum:
relevant at low pressures
Second continuum:
relevant at pressure >100 mbar
Gas now transparent (no rare gas molecules around)
Vibrationally excited dimer
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 18: Scintillation
Valid for all the rare gases
M. Hofmann (2012) PhD dissertation, Technical University of Munich (and references)
Often wavelength shifter are needed solid and/or gaseous
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 19: Quencher and shifter
Quencher/shifter
Quencher/shifter
Another form of quenching:
absorption (not transparent admixtures).
E.g., adding a small amount of xenon into argon
results in xenon scintillation.
Energy transfer (favoured) from argon to xenon.
Then the xenon chain is followed.
With the addition of molecules
Concurrent mechanism somewhere in the chain:
quencher without and shifter with radiative emission.
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 20: CF4
Admixtures which provide useful scintillation
(enough and near the visible range) are rare.
CF4 is not transparent to the CF4 de-excitation to ground.
What scintillates must be something else: (CF3*), CF3+* and CF4+*
A. Morozov et al., NIM A 628 (2011) 360
Pure CF4 primary scintillation
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 21: Where to get the light from
Primary
Excitation and ionisation produced by the particle interacting with the gas.
Some dependance
with the electric field.
Electroluminescence
Scintillation without charge amplification.
Increase linearly with the field.
Easy in pure noble gases.
Avalanche
Exponential scaling with the field.
Proportional to the charge gain.
The extreme is the visible spark.
Almost the same spectra
If only CF3+* and CF4+* are the responsible of the scintillation,
electroluminescence does not take place.
Certainly CF3+* and CF4+* are produced in avalanches,
in fact avalanche scintillation present in CF4 mixtures.
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 22: Overview
Relevant mechanisms for the operation of gaseous detectors
Imaging applications of MPGDs outside high energy physics
Developments of the optical readout for MPGDs at CERN
Possible other applications
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 23: Imaging
Visual representation of an object (often, but not necessarily in spatial coordinates) using any of a variety of techniques.
When using a camera it’s called photography.
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 24: Why imaging?
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 25: Sight
Our dominant sense:
- Intuitive handling of the data
- Immediately available information
- Visual learning fast and effective
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 26: Radiography with GEMs
https://gdd.web.cern.ch/GDD/gemreadout.htm
Demonstration application:
first radiography with GEMs of a small mammal.
Cu X-ray and triple-GEM with X-Y readout.
11 mm
6 cm
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 27: C-RAD GEMINI
gamma
from 6 MV linac
Increasing demand for better accuracy in monitoring the tumour position and movements.
Online verification of the actually delivered dose distribution.
Detectors exposed to high dose: radiation hard and ageing matters.
GEM
pixelated readout
dedicated
electronics
patient
+
-
+
-
converter
electron
Clinical installation
40x40 cm2
Imaging applications in the radiotherapy treatment process
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 28: C-RAD GEMINI
Pelvic phantom
QA object (resolution and contrast measure)
40 cm
6 MV linac, 1.5 MeV average gamma energy, 3 MU dose
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 29: X-ray fluorescence
A. Zielińska, 2013 JINST 8 P10011
Mock painting done with Renaissance techniques
Scanning XRF revealed a painting under Rembrandt’s “An old man in military costume”.
It took 19 days with an X-ray tube and 24 h with synchrotron radiation.
Scanning is detailed, but long and expensive.
Position and energy sensitive method
No position scan required
The system can be brought to the museum
XRF imaging system for fast mapping of pigment distributions in cultural heritage paintings
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 30: X-ray fluorescence
A. Zielińska, 2013 JINST 8 P10011
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 31: Overview
Relevant mechanisms for the operation of gaseous detectors
Imaging applications of MPGDs outside high energy physics
Developments of the optical readout for MPGDs at CERN
Possible other applications
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 32: Optical readout
Record the light emitted during the Townsend avalanche with a camera: use the detector as a scintillating plate.
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 33: Only techniques are new
G. Charpak et al., NIM A258 (1987) 177
Parallel mesh chamber filled with Ar/CH4/TEA 80%/8%/2%
seen by an image intensifier and a camera
Muons and delta rays
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 34: Advantages
Simplicity: like taking a picture
Robustness: as a device off-the-shelf
Versatility: several uses / environments
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 35: CCD vs CMOS
Traditionally and due to manufacturing processes
Anyway, sCMOS very attractive (expensive)
CCD
low noise
~fast
sensitive
high power
many pixel
expensive
CMOS
higher noise
fast
less sensitive (not full frame)
low power
less pixel (partially true)
~cheaper
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 36: CCD
ICCD
Photocathode
MCP (typically)
Phosphor screen
Focus (maybe)
CCD
EMCCD
e- amplification in Si (impact ionisation)
Similar to avalanche photo-diode
Amplification after the CCD
Amplification before the digitisation
Reduce the importance of read-noise
Introduce the amplification noise
ICCD
EMCCD
Full frame / Interline
Back-illuminated
Noise (cooled, etc…)
Colours, micro-lenses
Size, aspect ratio
ADC depth, …
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017
From somewhere in internet

## Slide 37: CCD
ICCD
EMCCD
Full frame / Interline
Back-illuminated
Noise (cooled, etc…)
Colours, micro-lenses
Size, aspect ratio
ADC depth, …
Full-frame
sensitive (better red)
possible back-illuminated
slow
shutter needed
possible anti-blooming
expensive
Interine
~sensitive (better blue)
no back-illuminated
fast
no shutter needed
typically anti-blooming
cheaper
Back-illuminated
Increase the quantum efficiency
Extend sensitivity to UV (~>200nm)
Expensive
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017
From somewhere in internet

## Slide 38: Noise sources
Dark current: √(In te / qe) and T dependent
Hot pixels: anomalous large In
Pixel non-uniformity: noise and QE
Read noise: both random and coherent
Shot noise: photo-statistic
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 39: Lens
Focal length and aperture
Minimum focussing distance
Field of view at MOD
Vignetting: sensor size
Lens resolution: pixel size and density
Magnification: M ~ f / (q - f)
Light collection:
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 40: Our camera and lens
QImaging Retiga R6
CCD: 2688x2200 4.54x4.54 um2 pixels
ADC: 14 bit
rate: 6.9 fps (20fps with binning)
read noise: 5.7 e- RMS
dark current: 0.0002 e-/p/s @ -20o C
trigger: external bulb + others
Navitar
focal length: 25 mm
aperture: f/0.95
Mount: C-Mount
Sensor type: 1’’ format
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 41: The setup
10 cm
M = sensor size / image size
M ~ 0.1, Ω ~ 5x10-4
This implies:
- large sensor - low noise
- fast lens
- a lot of light
Geometrical acceptance
10 mm
2 mm
2 mm
2 mm
Ar/CF4 80/20
Plexiglass
+3 dioptre lens
25 mm f/0.95 lens
CCD camera
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 42: Ar/CF4 80/20 scintillation
About 0.4 ph/e- in the visible
Plexiglass cuts below 400 nm
Nicely overlapping with CCD efficiency curve
Preliminary
CF4
ArI 4p 4s lines
High charge gain -> high light gain
Light yield comparison
NaI: ~40 ph/keV GEM at a gain of 105: ~106 ph/keV
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 43
4.4 cm
55Fe X-ray imaging
One of the first images
20 min exposure
70 um hole, 140 um pitch

## Slide 44: X-ray images
Raw data: fast (<1 s) acquisition and no processing time
Charge acquisition (26/10/1998)
~7 cm
Light acquisition
X-ray tube with W target at 20 kV - 40 kV at few mA
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 45: X-ray images
X-ray tube with W target
11 kV
40 kV
Straight out from the camera, no even the flat field correction
8 cm
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 46: X-ray images
Increasing energy:
- X-ray more penetrating
- Worse position resolution (larger charge cloud)
X-ray interacting with the CCD
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 47: Fluoroscopy
50 ms exposure
10 Hz acquisition
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 48: CT and 3D imaging
Image -> Sinograms -> Filtered Back Projection -> 3D image
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 49: Fluorescence
X-ray tube
W target
40 kV, 1 mA
Contemporary industrial painting
Pin-hole
GEMs
Shielding
Camera and lens
17 cm
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 50: What should be expected
W target:
bremsstrahlung +
characteristic lines
Lα Lβ Lγ
Limiting factors (because not optimised):
- low energy -> window and cathode
- high energy -> gas
feasible:
Cr, Fe, Ni, Cu
difficult:
Au, Pb
more difficult:
Al, Ag
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 51: The painting (visible)
Fe
SS
Au
Al
?
Cu
Cu
Ag
17 cm
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 52: The painting (X-rays)
O(1 kHz) in 500 s exposure
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 53: Event by event: alphas
Settable GEM gain: increasing it
220Rn & 216Po
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 54: Event by event: electrons
Electrons from 125 keV γ-rays from 55Fe source
Settable GEM gain: increasing it more
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 55: Other examples
α from 220Rn & 216Po
X-rays from 55Fe
muon and δ-ray
shower from a charged pion
For hours in sealed mode at test beam
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 56: X-rays in details
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 57: Analysis of single clusters
Spatial resolution in the limit of long exposure:
280 um for 1 cm drift and 5.9 keV X-rays.
Dominated by e- diffusion and cluster size.
Average profile of 5.9 keV events
Resolution improves using the centre-of-mass of the cluster
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 58: The spectrum
from CCD images
FWHM at 5.9 keV over 8x8 cm2
~36% before gain uniformity correction
24.7% after gain uniformity correction
Full energy
Pileup
Escape
55Fe source
Gain uniformity, vignetting plays a role too
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 59: Fluorescence event by event
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 60: Fluorescence event by event
some color coding
Red
Green
Blue
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 61: Fluorescence event by event
some color coding
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 62: Fluorescence event by event
It’s Zinc
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 63: Crystallography
Detection efficiency -> Thick conversion gap
Parallel field: parallax error
Parallax error mitigation with radial field
It would improve the spatial resolution
of X-ray fluorescence setup too
field shaper
segmented GEM
segmented cathode
drift
radius
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 64: Flat exposure
Parallel field
Radial field
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 65: Radial exposure
Mask
Mask
10 cm
Parallel field
Radial field
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 66: Overview
Relevant mechanisms for the operation of gaseous detectors
Imaging applications of MPGDs outside high energy physics
Developments of the optical readout for MPGDs at CERN
Possible other applications
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 67: Beam monitor
Though to cope with 108 - 109 particles/s/mm2 -> Single GEM
More GEM are an option for weaker beams -> more material budget
beam
In principle very low material budget possible < 200 um H2O eq
Beam position and shape measurement (at 20 Hz or more)
For instance for ion therapy beamlines
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 68
0 s
0.5 s exposure
0.5 s
1 s
1.5 s
2 s
2.5 s
3 s
3.5 s
4 s
~10 cm

## Slide 69: Integrating
Intensity profile
Pion beam of about about 2 x 106 pions per spill
(two orders of magnitude less than the expected conditions)
Detector operated in sealed mode though not though for this.
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 70: Other kind of MPGDs?
H. Takahashi, “Development of a transparent Single-grid-type MSGC based on LCD technology,” MPGD2015
T. Fujiwara, “Development and application of Scintillating Glass-GEM detector,” MPGD2015
Scintillating glass THGEM
Transparent MSGC based on LCD technology
Developments not limited to triple GEMs
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017
See also other nice developments of optically readout triple GEM
M. Marafini et al., Nucl. Instrum. Methods A824 (2016) 562.

## Slide 71: TPC readout
Direction-Sensitive Dark Matter Searches with a low pressure CF4 detector with optical readout.
CERN/GDD development for studying scintillation in controlled conditions targeted to TPC readout
with emphasis also on the primary scintillation.
nuclear recoil
104 keVr (~66 keVee)
DRIFT
N. S. Phan et al.,
arXiv:1510.02170
Triple GEM in 100 Torr CF4
C. Deaconu et al., Physics Procedia 61 (2015) 39
DMTPC
Grids in 50 Torr CF4
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017
F. Brunbauer et al.,
Symposium on large TPCs (2016) Paris

## Slide 72: Possible other applications
UV imaging
Neutron imaging
Gamma imaging
Portable devices
Educational purposes
…
Seminar at SPhN department - CEA/Saclay - Friday 3rd of February 2017

## Slide 73
Freeze-frame of an X-ray movie of a flying drone
Radiography of a bat and closeup of the GEM holes
Radiography of a crushed cup with pens and its 3D tomographic reconstruction
Visible picture of a painting and its X-ray fluorescence image. Different colours refer to different materials (energy resolved)
Cu
Au
Ag
Zn
SS
Fe
Cu
1€
Al
Single X-rays from 55Fe and the energy spectrum extracted from the images
Alphas from 220Rn decay and its daughter 216Po
O (100 keV) electrons ~MIPs